## Zedi Calibration Form
Field Calibration Certificate
Customer:
Legal Site Description:
|  | Service/Work Order #: |  |  |  | Ambient Temp. |  | deg C | Suite 500, 600 - 6th Avenue SW Calgary, AB T2P 0S5 Tel: (403) 444-1100 Fax:(403) 444-1101
|  | Device Type: | Smart-Alek® |  | S/N:
|  | Pressure Core Ratings (URL): | dP | 125 | SP | 1450
|  |  |  | 2 |  | 4 | [7 empty] | Uncertainty Settings (1st row is default)
| dP % | iwc | iwc | kPa | psi | kPa | SP % | Device |  |  |  | dP units | SP units | dP as Found Default | dP as Left Default | SP as Found Default | SP as Left Default
| 0.2% | 24 | 24 | 6 | 72.5 | 500 | 0.2% | 2 |  |  |  | 2 | 2 | 0.2% | 0.2% | 0.2% | 0.2%
| 0.2% | 125 | 125 | 32 | 435 | 3000 | 0.2% | [6 empty] | Alternate | Alternate | Alternate | Alternate
| 0.2% | 300 | 300 | 74.6 | 1000 | 6895 | 0.2% | Smart-Alek® |  |  |  | kPa | kPa | 0.1% | 0.1% | 0.1% | 0.1%
| 0.2% | 500 | 500 | 124 | 1450 | 10000 | 0.2% | EFM Walk-up™ |  |  |  | iwc | psi | 0.2% | 0.2% | 0.2% | 0.2%
| 0.2% | 522 | 520 | 130 | 2500 | 17240 | 0.25% | [6 empty] | 0.25% | 0.25% | 0.25% | 0.25%
D I F F E R E N T I A L P R E S S U R E
Calibrator Mfr & Model # | [7 empty] | Serial No.
Calibrator Setting/Range | [7 empty] | Calibrator Calibration Due Date
Readings as found * |  |  |  | Calibration adjustment if necessary |  |  |  | Readings after calibration
(Verify Menu or RTM ) |  |  |  | (Calibration Menu ) |  |  |  | (Verify Menu or RTM )
| dP range: 0 to |  | iwc |  | dP range: 0 to |  | iwc |  | dP range: 0 to | 125 | iwc
% of full scale | Target (iwc) | Actual Rdg (iwc) |  | % of full scale | Target | DP Adjustment Value |  | % of full scale | Target (iwc) | Actual Rdg (iwc)
0 | 0 |  |  | 0 | 0 |  |  | 0 | 0
50 | 62.5 |  |  | 100 | 125 |  |  | 50 | 62.5
100 | 125 |  |  | Field Notes: |  |  |  | 100 | 125
80 | 100 | [6 empty] | 80 | 100
20 | 25 | [6 empty] | 20 | 25
0 | 0 | [6 empty] | 0 | 0
Limits*: | 0.2% | 0.25 | iwc | * Limits are at calibration/reference ambient temp. |  |  |  | Limits*: | 0.2% | 0.25 | iwc
S T A T I C P R E S S U R E
Calibrator Mfr & Model # | [7 empty] | Serial No.
Calibrator Setting/Range | [7 empty] | Calibrator Calibration Due Date
Readings as found |  |  |  | Calibration adjustment if necessary |  |  |  | Readings after calibration
(Verify Menu or RTM ) |  |  |  | (Calibration Menu ) |  |  |  | (Verify Menu or RTM )
| SP range: 0 to |  | psi |  | SP range: 0 to |  | psi |  | SP range: 0 to | 1450 | psi
% of full scale | Target (psi) | Actual Rdg (psi) |  | % of full scale | Target | SP Adjustment Value |  | % of full scale | Target (psi) | Actual Rdg (psi)
0 | 0 |  |  | 0 | 0 |  |  | 0 | 0
50 | 725 |  |  | 100 | 1450 |  |  | 50 | 725
100 | 1450 |  |  | Field Notes: |  |  |  | 100 | 1450
80 | 1160 | [6 empty] | 80 | 1160
20 | 290 | [6 empty] | 20 | 290
0 | 0 | [6 empty] | 0 | 0
Limits*: | 0.2% | 2.9 | psi | * Limits are at calibration/reference ambient temp. |  |  |  | Limits*: | 0.2% | 2.9 | psi
| Differential Pressure is: |  | iwc | SP effect on dP: If known, use normal operating static pressure; <----- otherwise use Full Scale pressure. -----> |  |  |  |  | Differential Pressure is: |  | iwc
| @ Static Pressure of: |  | psi | [5 empty] | @ Static Pressure of: |  | psi
T E M P E R A T U R E (RTD)
Temp. Standard Mfr & Model #
Temp. Standard Serial No. | [7 empty] | Standard Calibration Due Date
Readings as found |  |  |  | Calibrate if necessary then verify ► |  |  |  | Readings after calibration
(Verify Menu or RTM ) |  |  |  | (Calibration Menu ) |  |  |  | (Verify Menu or RTM )
Target Temperature | Test Std Value (°C) | Actual Rdg (°C) | Limit +/- | For typical well, use limits as listed (i.e., per regulators such as ERCB, BCOGC, & BLM). For API 21.1 compliance or cross-border measurement use ±0.25 °C (±0.5 °F) |  |  |  | Target Temperature | Test Std Value (°C) | Actual Rdg (°C) | Limit +/-
Above Process |  |  | 1.0 °C |  |  |  |  | Above Process |  |  | 1.0 °C
Process Temp. | [7 empty] | Process Temp.
0.0 °C | [7 empty] | 0.0 °C
Configuration for verification readings (as found and as left) must include the RTD probe connected to the device
1
Celcius | 1.0 °C | 0.0 °C | (°C) | [5 empty] | Pass Button Status | FALSE
Fahrenheit | 1.8 °F | 32.0 °F | (°F) | [5 empty] | Fail Button Status | FALSE
| Upstream Meter Run |  |  |  | Downstream Meter Run |  |  |  | Orifice Plate/Throat
Add any additional customer requested comments, plate size, condition, if meter type is eTube, etc. Please note bores for D, d, and ref temp. List units.
Performed by: | [5 empty] | Date: |  |  | Pass Fail
CONTRACTORS - Please leave original Zedi tag on unit, and add your new calibration tag indicating calibration status and date completed.
FM-CSS-05 Field Calibration Certificate | [10 empty] | Revised Dec 2008
Note on G4: Ambient Temperature must be listed, particularly when a Calibration Adjustment is made.
Print
Sign

## Revision History
Release: | Description of Changes:
2008-12 | Adjusted by Brian Z.
• Now allows some leeway in EFM Walkup Full Scale settings (+1%/-5%).
• Adjusted 520 iwc Full Scale for EFM Walkup to show 522 iwc (130kPa).
2008-11 | Adjusted by Brian Z.
• Moved the Date to the bottom with the signature.
• Added a reference cell for Service Call / Work Order number.
2008-09 | Adjusted by Brian Z.
• Can select whether it is a Smart-Alek or an EFM Walk-up
• The Full Scale span value can be adjusted to a lower value
• Can separately select Units of Measure for dP and for SP
• Added a Celcius/Fahrenheit selector
• Added a couple of notes areas in the dP and SP sections
• Separated out the signature section for the printed name and signature
• Listed the limits specification for pressures. i.e., added the % value
• Changed (to a Corporate) colour scheme
• Added validation to Span setting cells
• If FS setting is inadvertently set as > the core value, core value is used in the calibration sections
• Added ability to set the % FS uncertainty level for As Found and As Left pressure core readings.
2008-02
DRAFT 1.0 ADD REV LEVEL TO FORM to css 05 and new date REVISED BY JEFF R. FEB 14 2008
CORRECT LOGO to corporate standard and confirm cgy address
Line 62 - Add us ds and orifice plate
Lines 56 to 58 - Temp targets changed to reflect zero shot, process shot, and third point above process or zero to verify
Brian Z will add revised comments regarding ERCB BCOG API temp standards shortly and .25 may change to 1 degc
lines 20, 36, 54 -- menus or RTM notes changed from RTM only notes
Line 59 -- temp hot keys opened up for degc and degf - where previously locked on degc
MACROS AND HOT KEYS ARE STILL PLACE
DRAFT 2.0 with Brian Z. requested changes attached Revised Feb 14 2008
E56 - Tried a new wording. (FYI - Strictly speaking BLM is +/- 2F, but 1.8F (1C) is close enough).
Quick notes:
- For readability used 0.25 instead of just .25.
- D40, Change out AEUB
- A68, changed Zedi instead of zed.i
- D40 delete ERCB and API reference as three point or 6 point. Default is now to test all points up and down in span
Feb 20 2008 appears acceptable per Tory Nudd, and Rod Nisi at Taqa North
Added revision date and form name. Footer deleted as it overwrrites cells on new form
2006-03 | • Updated corporate address
• Added clarifications for RTD verification points
2005-06 | Updated and released with Form number FM-CSS-05.
• Added input boxes for calibration equipment info, Ambient Temperature, etc.
• Added limits values
• Added section for Line Pressure effects on dP
2002 | Initial Release as "Smart-Alek Calibration Certificate" by Len Benson, reviewed by Brian Z.